## GAS FLOW
CIPET:IPT:AHMEDABAD, GOVERNMENT OF INDIA
Central Institute of Petrochemical Engineering & Technology
(Department of Chemicals & Petrochemicals, Ministry of Chemicals & Fertilizers)
Plot No.630,Phase-IV,G.I.D.C.,Vatva,Ahmedabad-382445.Gujarat.INDIA.
Technical Bid for GAS FLOW CALIBRATOR
TENDER NO. | CIPET:IPT/AHD/CALIBRATION/2021-22/05
ITEM DESCRIPTION | PROCUREMENT OF LABORATORY EQUIPMENTS.1).GAS FLOW CALIBRATOR,2).AUTO TITRATOR,3).MULTI FUNCTIONAL CALIBRATOR,4).MODA.
TECHNICAL SPECIFICATION OF GAS FLOW CALIBRATOR - 01 Nos.
NAME OF THE SUPPLIER / MANUFACTURER
Sl. No. | PARTICULARS |  | REMARKS
1 | Flow range | 0.1 to 10 slm (standard liter per minute)
2 | Uncertainty | 0.5% reading
3 | Precision | 0.025% reading
4 | Stability | 0.15% reading
5 | Operating temperature | 15 °C to 30 °C
6 | Communication Interface | RS 232, IEEE
7 | Power Requirements | 85 V ac to 264 V ac, 47 Hz to 440 Hz.
8 | Control | Complete front panel control
9 | Advanced functions | totalize, average, hi/lo, deviation, on-board purge, leak test and tare
10 | General Specification | Real time True mass flow calibration
11 |  | Both mass and volume flow with user settable reference pressure and temperature conditions
12 |  | Gas flow regulation and adjustment hardware. Modular component allows future upgrade and expansion in range
13 |  | Supports 14 different gas calibration
14 |  | Includes traceable calibration for operation in N2 with corrections for other gases
15 |  | Internal valving for on-board purge, leak test and tare support

## MULTIFUNCTION
CIPET:IPT:AHMEDABAD, GOVERNMENT OF INDIA
Central Institute of Petrochemical Engineering & Technology
(Department of Chemicals & Petrochemicals, Ministry of Chemicals & Fertilizers)
Plot No.630,Phase-IV,G.I.D.C.,Vatva,Ahmedabad-382445.Gujarat.INDIA.
Technical Bid for MULTIFUNCTIONAL CALIBRATOR
TENDER NO. | CIPET:IPT/AHD/CALIBRATION/2021-22/05
ITEM DESCRIPTION | PROCUREMENT OF LABORATORY EQUIPMENTS.1).GAS FLOW CALIBRATOR,2).AUTO TITRATOR,3).MULTI FUNCTIONAL CALIBRATOR,4).MODA.
TECHNICAL SPECIFICATION OF MULTIFUNCTIONAL CALIBRATORS- 01 Nos.
NAME OF THE SUPPLIER / MANUFACTURER
Sl. No. | PARTICULARS |  | REMARKS
1 | DC Voltage Measurement | Range – 0 to 55 V
Resolution-0.001 V
Accuracy-0.0015 %
2 | DC Voltage Source | Range – 0 to 33 V
Resolution- 30V-0.001 V, 100mV- 0.01 mV
Accuracy-0.0015 %
3 | DC mA Measurement and Source | Range – 24 mA
Resolution- 0.001 mA
Accuracy-0.01 %
4 | MilivoltMeasurement and Source | Range – -10 mV to +75 mV
Resolution- 0.01 mV
Accuracy-0.02 %
5 | Ohms Measurement/ Source | Range – 0 to 4.4kΩ
Resolution- 1 mΩ
Accuracy-0.015 %
6 | Frequency Measurement/Source | Range – 0 to 50 kHz and CPM mode for pulse
Resolution- 0.01 Hz
Accuracy-0.005 %
Thermocouples
Type | Range ( °C) | Resolution | Accurency
J | -200 to 1200 | 0.1°C | 0.8°C
K | -200 to 1370 |  | 0.5°C
T | -200 to 400 |  | 0.5°C
E | -200 to 950 |  | 0.5°C
R | -20 to 1750 |  | 0.8°C
S | -20 to 1750 |  | 0.8°C
B | 600 to 1800 |  | 0.8°C
L | -200 to 900 |  | 0.5°C
U | -200 to 400 |  | 0.5°C
N* | -200 to 1300 |  | 0.8°C
XK | -200 to 800 |  | 0.8°C
BP | 0 to 2500 |  | 0.8°C
RTD
Type | Range ( °C) | Resolution | Accurency
Ni120 | -80 to 260 | 0.1 ° C | 0.15°C
Pt100-385 | - 200 to 800 |  | 0.45°C
Pt100-392 | -200 to 630 |  | 0.35°C
Pt100-JIS | -200 to 630 |  | 0.35°C
Pt200-385 | -200 to630 |  | 0.95°C
Pt500-385 | -200 to630 |  | 0.55°C
RAMP FUNCTION
Type
Source functions | Voltage, current, resistance, frequency, temperature
Ramps | Slow ramp, Fast ramp, 25% step-ramp
LOOP POWER FUNCTION
Type
Voltage | 24 V
Accuracy | 10%
Maximum current | 22 mA, short circuit protected
STEP FUNCTION
Type
Source functions | Voltage, current, resistance, frequency, temperature
Steps | 25% of range, 100% of range
| OPERATING TEMPRATURE | -10°C TO 55°C -10°C TO 55°C
| STORAGE TEMPRATURE | -20°C TO 70°C (limited by battry storage specification)
| Stability | ±0.005% of range °C outside of 23± 5°C
| Operatingaltitude | 3000 meters above mean sea leavel
| Safely and EMC | IEC 61010-1, IEC 61010-2-030, Pollution Degree 2 IEC 61326-1 (Portable), CISPR 11 (Class A)

## AUTO TITRATOR
CIPET:IPT:AHMEDABAD, GOVERNMENT OF INDIA
Central Institute of Petrochemical Engineering & Technology
(Department of Chemicals & Petrochemicals, Ministry of Chemicals & Fertilizers)
Plot No.630,Phase-IV,G.I.D.C.,Vatva,Ahmedabad-382445.Gujarat.INDIA.
Technical Bid for AUTO TITRATOR
TENDER NO. | CIPET:IPT/AHD/CALIBRATION/2021-22/05
ITEM DESCRIPTION | PROCUREMENT OF LABORATORY EQUIPMENTS.1).GAS FLOW CALIBRATOR,2).AUTO TITRATOR,3).MULTI FUNCTIONAL CALIBRATOR,4).MODA.
TECHNICAL SPECIFICATION OF AUTO TITRATOR - 01 Nos.
NAME OF THE SUPPLIER / MANUFACTURER
Sl. No. | PARTICULARS |  | REMARKS
1 | Milli Volt | ±2500mv
2 | Accuracy | ±0.01mV
3 | Input Impedance | 1012Ohms
4 | Burette Resolution | 1 ul ( with 20ml Burette) Interchagelable
5 | Filling Time | ˂20Sec.
6 | Sensors | Ph,Redox,Aqueoous,Non-aqueous,Complex metric,Argentometric,ISE,Temperature
7 | Power | 230V,50Hz
8 | Stirrer | Separate Stirrer with Touch Screen control
9 | Display | 7” Color Capacitive Touch Screen
10 | Calibration | 5 point calibration
11 | Report & Method Storage | 1000 result sand Method Storage
12 | Technology | ARM Cortex SOC with MIPS
13 | Cut-off Criteria | Volume,End Point,Mv/Ph
14 | Communication Interface | Modbus( Intergrated software for multiple instruments,Data collection with 21 CFR Part II Compliance),Ethernet/WIFI(PC connectivity,printer connection over intra-network) USB Port for Data Transfer,Printer Interface
15 | End Point Detection | Volumetric,Potentimetric
16 | Reports | a.Analysis Report: Result,Parameter,Graphs
17 |  | b.Statistic Report: RSD,Average,Linearity
18 | Units | Conc,%W/W,%W/V,ppm.gm/L,mg/L,TAN/TBN
19 | Titration Mode | Blank,Sample,Normality,Titer( for KF Titrator)
20 | Methods | SVD,ASD,FED & KF( with optional attachment)
21 | Ep Detection Criteria, 0.8°C | Largest,Last & Selected EP
22 | Environmental Condition, 0.5° | Operation: indoor,Temperature: Ambient to 45° Humidity:5 to 90% non Condensing.

## MODA
CIPET:IPT:AHMEDABAD, GOVERNMENT OF INDIA
Central Institute of Petrochemical Engineering & Technology
(Department of Chemicals & Petrochemicals, Ministry of Chemicals & Fertilizers)
Plot No.630,Phase-IV,G.I.D.C.,Vatva,Ahmedabad-382445.Gujarat.INDIA.
Technical Bid for MODA (Microbial Oxidative Dergadation Analyser)
TENDER NO. | CIPET:IPT/AHD/CALIBRATION/2021-22/05
ITEM DESCRIPTION | PROCUREMENT OF LABORATORY EQUIPMENTS.1).GAS FLOW CALIBRATOR,2).AUTO TITRATOR,3).MULTI FUNCTIONAL CALIBRATOR,4).MODA.
TECHNICAL SPECIFICATION OF MODA-01 Nos.
NAME OF THE SUPPLIER / MANUFACTURER
Sl. No. | PARTICULARS |  | REMARKS
A | Data is accurate and reliable
1 | The high-precision gas mass flow controller ensures the stability and accuracy of the intake air flow.
2 | Imported carbon dioxide sensor to ensure the stability and high precision of the measurement results.
3 | Built-in cavity measurement calibration system, convenient to detect the accuracy of carbon dioxide and oxygen.
Easy to operate
4 | Professional software support, the software interface is simple and clear, easy to operate, and the test process can be set flexibly.
5 | Fully automatic operation, setting the corresponding parameters, one-key testing.
6 | The software displays and records the experimental status and process parameters in real time.
7 | Professional test report generation, can export files in Office, PDF and other formats.
Technology advancement
8 | Built-in automatic mixing system, can customize mixing cycle and mixing time.
9 | PID fuzzy algorithm temperature control technology to achieve high-precision temperature control.
10 | Built-in leak detection system to ensure the air tightness of the gas path.
High test efficiency
11 | Eighteen cavity samples are independently tested, and each cavity is individually data collected.
Authority management and data tracking
11 | A user name and password are required to log in to the workstation.
12 | Users are divided into multiple levels of operators and administrators (such as administrators, operators, observers, etc. but not limited to these levels).
13 | The administrator can adjust the authority of various levels; for example, increase and decrease system control items of a certain level.
14 | With audit trail function (system audit trail, project operation audit trail, method audit trail), every data change is recorded; the safety and integrity of test data are guaranteed.
The instrument is stable, reliable and easy to maintain
15 | Professional carbon dioxide, oxygen, pressure sensor, high precision, good stability, can run for a long time
16 | Functional modular design, easy to maintain.
Humanized design
17 | The test process can be viewed on the computer and mobile phones,and the experimental data can be monitored in real time
18 | Equipment networking status, which can provide online control, upgrade and remote service functions.
19 | Direct measurement of CO2 for the Bio-degradable Plastic degradation by Micro organism.
20 | Should comply IS/ISO: 17088,ISO:14885:Part-II,ISO:13975 and ASTM:5338
TECHNICAL SPECIFICATION
Sl. No. | DESCRIPTION | RANGE
1 | Temperature Control | Room Temperature ~80℃
2 | Temperature Control Accuracy | ±0.5℃
3 | Flow Control | 0.01~10L/min
4 | Flow control | rate ±1.0%FS
5 | Flow Detection | 0.01~10L/min
6 | Flow Detection Accuracy | ±1.0%FS
7 | Carbon Dioxide Detection | 0~10000ppm
8 | Carbon Dioxide Detection Accuracy | ±1%FS
9 | Oxygen Detection | 0~25%
10 | Oxygen Detection Accuracy | ±1%FS
11 | Test Chambers | 18 chambers
12 | Pressure Detection | 0~200KPa
13 | Pressure Detection Accuracy | ±1%FS
14 | Air Pressure | 0.1 ～ 0.5Mpa
15 | Air Connection | 1/4 Inch Hose
16 | Test Cavity Volume | 3L
17 | Power | 3000W
18 | Power Supply | AC 220V，50Hz