## Sheet1
Technical Specification for Tender No.2022-23/02 - Released in Government-e-Marketting (GeM) Portal
3D Printer
Sl. No. | Specification | Range / Value
| Make | Bidder to specify
| Model | Bidder to specify
1 | MACHINE
1.1 | Technology | Solid based additive manufacturing system (Latest Fused Deposition Modelling system)
1.2 | Build Volume | Minimum build volume 350 * 350 * 350 mm (±50mm)
1.3 | Nozzles | 2 (Dual Wear Resistance Nozzles) for model and support materials.
Nozzles for layer thickness: 0.3, 0.4, 0.6 mm
1.4 | Layer resolution | Minimum 20 micron or better
1.5 | Part accuracy | Parts produced should be within an accuracy of +/- 0.15 mm
1.6 | Support structure | Soluble and breakaway
1.7 | Machine Control | Machine should build parts in circulated air heated chamber to provide uniform temperature control and heated build environment throughout the volume of the build and each layer of the print.
Chamber Temperature: as per the requirement of printing material
Non heated removable and flexible build plate/sheet made form heat absorbing material such as Polycarbonate or ceramic build plate or equivalent
Motion controlled components should place outside and should be insulated from the heated build chamber.
The support material distinct material which should extrude from a different printhead/nozzle.
The machine shall automatically calibrate nozzle and build platform in X, Y and Z axis before the start of each job.
Auto and manual calibration of offset between model and support nozzle.
The extruder Nozzle tips should be self cleaning
1.8 | Filament Flow Control | 3D Printer sholuld have Automatic Filament Flow Control System for easy and smooth dispensing of material
1.9 | Air Circulation System | 3D Printer should have air circulation system
1.1 | Material Flow Rate | 50 mm3/sec or better
1.11 | Extruder/ Nozzle Temperature | 450 °C or Better
1.12 | Monitoring | Live camera for tracking and monitoring, mobile app for android and iOS should be available
1.13 | Material Handling | Material handling systems should be part of the Printer with automatic material loading, feeding and storage management system. Material cabinet should be dry sealed with humidity control. All material spool/cartridges (model and support) should have an EEPROM /chip to indicate the quantity of material available in the spool / cartridge at any instance of the machine operation during idle or run time. Tip life should be displayed on the monitor.
For large jobs. The machine shall automatically pause and prompt the operator to replenish material and resume printing in case of the material (for both model and support) gets consumed during the build process. The machine should keep the chamber temperature and heated airflow constant while the machine is paused.
1.14 | Connectivity | Wi-Fi, USB, LAN/SD card compatible
1.15 | User Interface | Touch Screen
Using computer with appropriate software. Software should display and control. Complete program of solution must be provided.
1.16 | Slicing software | Software should be able to import any native CAD formats like stl, obj igs, vrml. stp, Parasolid, native files of CAD packages etc and be able to repair errors like open surfaces, inverted normal etc without use of any 3rd party software.
Software should be from same OEM; 3rd party software is not allowed. License must be perpetual and cloud support must be available.
Software should allow by changing the parameters / add or deletion of support, infill percentage, etc.,
2 | Materials
2.1 | Materials | Standard Materials (Print Settings should be pre-configured in machine and software): ABS, ASA, Chemical Resistant Material, Nylon12, Carbon Fiber re-enforced material, PC-ABS, Fire-Resistant PC-ABS
Research Material: User should be able to run self-developed/open-source materials metal and polymer filaments.
Vendor should quote all possible materials and supply 5 Nos. of Filament Spool on each material, excluding the quantity of material required for installation, training, and calibration. Machine should be calibrated to fabricate parts with all the materials mentioned above at the time of installation.
Vendor should attach the quality certificate of each material with test report evaluating the parameters according to ASTM/DIN standards tested on same/similar machine from same OEM.
2.2 | Customised material guidance | Machine should have the provision of testing materials developed by the customer and professional guidance should be given to adhere to the system compatibility and configuration
2.3 | Material Quality and Property | Vendor should also supply proven parameters for all the suitable materials with clear documentation and statistics of the mechanical properties with respect to build orientation and layer thickness.
3 | ESSENTIAL ACCESSORIES
3.1 | Support removing system | As per the requirement to be quoted by the bidder (for soluble support removal)
3.2 | Consumables | Vendor should supply minimum quantities of consumables like build platforms, nozzle for all layer thickness, wiper blade assembly with purge ledge, brush etc., required for 6 months
3.3 | Tool kit | Vendor should supply standard tool kit for start-up, removal of parts and cleaning (list to be attached).
3.4 | Any other accessories required | Vendor should supply all the other accessories, material transport trolleys / carts and spares required for effective and better utilization of machine. All the required accessories should be listed
3.5 | Compressor | Vendor should supply Suitable compressor with dryer and all necessary accessories (if required)
3.6 | De-humidifier | Vendor should supply suitable de-humidifier to maintain room humidity level within suitable range for machine operation (if required).
3.7 | Online UPS | Vendor should supply suitable UPS with minimum 60 minutes power backup for the machine and essential accessories. Should have built in safety to protect machine from voltage spikes and sudden surges.
3.8 | Workstation with accessories | Vendor should supply suitable OEM workstation computer system with complete accessories and UPS. It should have minimum of i7 (3.4 GHz, quad core) processor of latest configuration, 64GB RAM (2400MHz, DDR4), 1TB hard disk (7200 RPM), and 4GB graphics card or better configuration, with windows operating system and MS office recent version.
4 | OTHER ESSENTIAL REQUIREMENTS
4.1 | Benchmark part | Vendor should submit two numbers of benchmark parts in two material which should be built in the quoted machine same is to be built during prove out of the machine at the site
4.2 | Manufacturer’s/Suppliers Credential Credentials | Supplier should have sizable installations of same model worldwide and at least three same or similar models in India. Supplier should submit complete contact details and acceptance letter from the recent customers.
Manufacturer of the quoted equipment should be ISO certified and should have supplied same/similar machine in India/globally for a period of minimum of 10 years.
4.3 | Safety Requirements | The machine and all the accessories supplied to meet objective should be able to operate without any risk or hazard, without any additional protection, provision, training or guarding devices and meet current international standards such as CE, FCC, EMC, Low Voltage Directive 60950-1, 62368 - Certificate of Conformance to be attached
4.4 | Scope of supply | Vendor should supply complete start up package necessary to prove the machine and provide training. List for scope of supply to be submitted.
4.5 | Environment Protection | The machine and all the accessories supplied should be safe to use without emission of any hazardous gases, noise level and radiation without any need for additional equipment, provision or training and meet current international standards
4.5 | Consumables & accessories and their availability | List all such material that will be used in building part. Tender shall include list of all essential spares and consumables to be provided with replacement time prescribed for each such item and its availability within reasonable time period. In case if any such item is likely to be out of availability within service period of machine, such item shall be included in initial supply
4.6 | Price list of material, spares, and consumables | Price list of each material with minimum order quantity, machine spares and consumables are to be quoted.
5 | INSTALLATION, COMMISSIONING AND TRAINING
5.1 | Installation and requirements | Vendor should state the space required and condition of floor and any other requirements for installation of the machine and equipment’s. State clearly the specifications of the following items required for installation such as Transformer, chiller, UPS, Air conditioner, Dehumidifier, separate special earthing, vibration isolation. Material disposal procedure for the soluble support system / agitation / cleaning tank
5.2 | Training and documentation | Minimum of 5 days training for five persons which includes basic & advanced level training. Training content and plan to be submitted. Training faculty must have adequate experience in this field.
The vendor should supply the necessary manuals such as
 Software instruction
 Maintenance and trouble manual
 Training
 Installation and Commissioning
 Handling of accessories
 Software key (if any)
 Software CDs
| Installation and requirements Commissioning | The Vendor should support necessary site preparation for installation. Vendor should cary out installation and commissioning of the machine and its accessories on a turnkey basis. Material disposal procedure for the soluble support system / agitation / celaning tank.
Colour Spectrophotometer
Sl. No. | Specification | Range / Value
1 | Make / Model | Bidder to specify
2 | Purposes | To measure Colour, Grey Scale rating and Yellowness Index as per various Indian and International standards.
3 | Applicable standard | ASTM D1925, ASTM E313,ASTM D2244, ASTM E308,ASTM E1164, DIN 5033, JIS Z 8722, ISO 7724/1, CIE 15: 2004
4 | Measurement principle& Mode | Dual beam spectrophotometer & Reflectance and Transmittance with UV cutoff filter
5 | Light Source | Pulsed Xenon lamp filtered to approximate D65 daylight
6 | Detector | Silicon Photodiode array
7 | Viewing Aperture | Large area view : 1.75 in illuminated, 50 mm measured
Small area view : 1.75 in, 1.00 in, 0.50 in, 0.25 in & 0.13 in illuminated
2.0 in., 1.20 in, 0.70 in, 0.40 in, 0.20 in, measured
8 | Lens switching for LAV/SAV | Automatic
9 | Spectral Range | 360 - 700 nm or better
10 | Resolution | <3nm
11 | Effective Band width | 10 nm equivalent triangular
12 | Photometric range | 0 - 150%
13 | Photometric Resolution | 0.01%
14 | Lamp life | 1 billion flashes or better
15 | Automatic UV Control | 420nm cut off filter for UV Control &UV Exclusion
16 | Measurement Time | < 3 seconds; (except 3mm area<10sec)
17 | Calorimetric repeatability | For white tile: ΔE* < 0.09 or better for 1.75 inch
For Blue denim tile: ΔE* < 0.07 or better for 1.75 inch
18 | Inter instrument agreement | ΔE*< 0.15 (Avg.) for 1.75 in.
ΔE*< 0.36(Max.) for 1.75 in. CIE Lab (max.)
19 | Equipment to be supplied with all essential Accessories such as: | Calibrated white UV Fluorescent Standard with NIST Traceable certificate of calibration - 01no
Sample Cup Opaque Cover - 01 no
Glass Sample Cup (2.5in) - 04 nos.
Port insert, 2.5in Glass sample cup holder-01no
Sample Clamp Assembly - 01 no
Other parts like cable, adopter, power cord, UPS etc.
Black Calibration light Trap
Green Check Tile
Operation manual
20 | Color matching software to measure: L, a, b, Δxyz | Suitable advanced software inbuilt with instrument to analyse L, a, b, ΔE, yellowness index, grey scale rating, etc. Software must be perpectual with all latest advance modules.
21 | Other Mandatory Accessories to be quoted and supplied along with machine / equipment | While supplying the Machine, the supplier should also provide the following items apart from above:
Branded PC of best configuration with necessary software including software for colour matching & suitable for the instrument operation, Colour printer
Basic tool Kit - 01 set
Hard copies of Operational & Service Manual - 01 set and certified reference material
Machine should come with all other essential accessories & spares (as per IS, ASTM & ISO standards) required for installation, commissioning & operation. Bidder should specify and quote for any other acessories required / availble for better usage of machine.
Onsite free operational Training
Calibration certificate traceable to NIST should be provided
22 | Scope of Supply / Bill of Material | Bidder should submit Scope of Supply / Bill of Material with make model of each items.
| Technical Support and Service | 1. The Supplier Shall be responsible for carrying out the installation and Commissioning at customer site. 2. Machine Should come with all other essential accessories & spares required for installation, commissioining & operation.
GC - MS With Pyrolyzer and Headspace
Sl. No | Item | Specification
1 | Applications | Separation and identification of solids/liquids/gaseous and their mixtures that includes unknown samples, polymers, pesticides, organic compounds, pharmaceutical ingredient etc.
2 | General Technical
| Gas Chromatograph system | Instrument detection limit should be ≤10 fg OFN for Statistically derived at 99 % confidence level from the area precision of 8 sequential splitless injections of 10 fg OFN standard MS/MS transition of m/z 272 & 222 using 30m long colum or similar
3 | Ionization mode | Electron Impact (EI)
4 | Detector | Flame Ionization Detector (FID) should be provided which can be coupled with the Gas chromatographic system.
5 | Data acquisition /operating system | Data acquisition /operating system System should be capable of supporting two inlets and TWO detector ports simultaneously and with Mass Spectroscopy; should have electronic pneumatic/pressure controls for all the gases and should have Chromatography Data system which is based on Microsoft Windows operating system for instrument control, data acquisition, data analysis, quantization, automation & customization with online and offline sessions provided.
6 | Column Oven
| Columns | Provision to install atleast two column
7 | Operating temperature | 450°C or more
8 | Temperature set point resolution | 1°C or better
9 | Maximum temp ramp rate | 120 °C /min or more
10 | Cooling rate | From 450 °C to 50 °C: within 4 minutes or better
11 | Temperature programming | Should have minimum 9 ramps & 10 Plateaus
12 | Head Space Analyzer
| Injection system | Pressure balanced system also along with loop based system
13 | Sample | Should able to handle all type of VOC
14 | Incubation Temperature | 35⁰C to 200⁰C or more
15 | Head space vial | Vial capacity 12 or more and upgradable to more vial capacity
16 | Pyrolyzer
17 | Make and model | Bidder to specify
18 | Type | pyroylyzer compatible with GC-MS
19 | Temperature range | Upto 1000 deg.C or better,Programmable Temperature,Auto sample positioning
20 | Furnace cooling rate | Temperature to go down from 800 deg.C to 50 deg.C within 10 min.
21 | Sample to be analyzed | Solid and viscous liquid
22 | Control | Should be provide with controll software
23 | Injection port
| Injection port-2 nos to quote | 1. Split/Splitles sinjection port with electronic pressure control (EPC)/ programmable pneumatic control(PPC) /advanced flow control (AFC) with fast GC capability
• Programmable vaporizer injector and programmable up to 2 ramps (or better); Heating rate 200 deg.C/min. or better
Possible to use capillary columns of 100μm to 530μm columns
Digital display of gas flow, temperature etc.
Manufacture’s software controlled (AFC/EFC/APC/EPC controlled).
24 | Pressure range | 100 psi or better
25 | Maximum temperature | 400 ⁰C or more
26 | Heating zones | Should have independently heated zones
27 | Auto Injector-liquid | An automatic injector device having a capacity to hold at least 15 vials capacity and more and should be field upgradabale to 150 vial capacity. Provision for inject both injectors without dismantling hardware or any manual alignment.
28 | GC Detector Specifications (FID)
| FID detector | Having an MDL:<3 pgc/s or better
| Linear dynamic range | 107 or better
| Mass range | m/z up to 1000 unit or better
| Mass Analyzer | Should have inert/metallic quadrupole massfilter with pre-filter or equivalent technlogy
| Mass axis stability | Should be ±0.10 amu over 48 hrs
| Scan speed | up to 10,000 u/sec or more
| Ion source temp | upto 350 ⁰C or better and it should be programmable.
It should be cable less source for easy cleaning and maintenance.
29 | Ionization mode | EI
30 | Filament | Minimum single filament
31 | The sensitivity of system should be a followed and demonstrated at site | EI 1500:1 for 1pg of OFN or better
32 | Turbo Molecular Pump (TMP) | 255 L/sec or better capacity
33 | Resolution | Selectable, 0.7 to 2.5 Daltons,
34 | WorkStation Instrument Control Software | GC and MS system should be combined with the same workstation for simultaneous settings and programming.
Should have Auto tune (to optimize MS parameters automatically) feature
A user friendly automatic data collection and analysis system compatible with Microsoft Windows OS and Microsoft Office suite applications
NIST Library should be provided.Support to be provided for updation of library
35 | Sample preparation | Qucheers Kit: 2 No.
36 | Spectral Library | Latest mass spectral library (NIST) to be supplied in CD (licensed) for polymers including rubbers, additives, pesticide, insecticide, etc.
37 | Accessories to be quoted and supplied along with machine / equipment | Branded latest suitable PC compatible with GC-MS system having HDD Graphics display, 20”LCD/LED Monitor alongwith a good quality printer (should specify the PC and printer model).
UPS: 10 KVA UPS with at least 60 min back up is needed.
Supporting Table for installation of equipment and computer to be provided
High quality He, H2, N2, Argon & Zero Air gas cylinders alongwith compatible regulators, gas purification panel for the above mentioned gases,and required tubings should be provided.
Syringes for manual and autosampler injection (minimum pack of 10 each), Filaments (minimum 2 no.), capillary column (minimum 1 no., desired 2 no.).
Mechanical accessories (tool-kit etc) and consumable spares (Vials, Septa, Ferrules etc.) for the operation and maintenance of the instrument should be provided to meet our needs for at least 1 year.
38 | Other terms and conditions | Installation Qualification (IQ) and Operational Qualification (OQ) should be performed at the time of installation and commissioning.
The system must be factory tested and a certificate should be provided
The entire system should be installed by the company professionals at oursite. A thorough technical training (minimum 3 days) in analyzing and troubleshooting should be given by the technical professionals
A list of references in India,where similar systems have been installed,must be provided and this will be taken very seriously while making the decision.Your post sales service feedback will certainly be a deciding factor
Complete set of manuals on operation, maintenance of the system in hardcopy as well as soft copy should be provided in English.
High Resolution Inductively Coupled Plasma Optical Emission Spectrometer (ICP - OES)
Sl. No | Item | Specification
1 | Introduction | A compact high performance Dual View ICP system is required for the routine measurement of major and minor elements in a variety of general samples and research samples.
The ICP should be true simultaneous using polychromator, in all aspects of measurements including simultaneous measurements of all analyte wavelengths, internal standard and background.
2 | Optical System | The instrument shall be equipped with a high energy Echelle Cross Dispersion Spectrometer with focal length of less than 400mm & a “side by side” arrangement of prism and grating.
Emission spectra shall be dispersed in both wavelength and order in a 2 dimensional array
The optical purge shall be with either argon or nitrogen to allow measurement of signals in the low UV part of the spectrum. The instrument shall operate with a reduced purge flow of 2 lpm and include a standby flow of 1 lpm to minimize gas consumption.
3 | RF Generator | The Solid state water cooled RF generator must run at a frequency of 40 / 27.12 MHz .
The generator shall have a high coupling efficiency of >75% and be directly coupled with a swing frequency impedance control and power regulation to better than 0.1%. RF power should be minimum 1500 watt in both axial and radial mode and variable with 5 watts or less in increament
4 | Plasma / Auxiliary / Nebulizer Gases | Argon Gas consumption should be as low as possible.
Preferably 8 to 10 lit/min or similar, including optics purge
Shear gas should be avoided for better resolution & sensitivity.
Variable flow rates for Nebulizer (0 to 2.0 L/min in 0.01 L/min increments), auxillary(0 to 2.0 L/min in 0.1 L/min increment) and plasma Argon flows for the better optimization of the method.
5 | Standard Sample Introduction Kit / Peristaltic Pump | Sample introduction shall be via an integral, close coupled and variable speed, peristaltic pump with a minimum of 12 rollers. The pump shall have a minimum of 4-channels to accommodate sample, drain, online internal standard and hydride generation line all together.
An high efficiency cyclone spray chamber fitted with a high performance glass concentric nebulizer shall be provided for optimum stability and best detection limits with fast washout and minimum cross contamination. Periataltic pump should be from 0.4 mL to 7 mL / min variable in 0.1 mL/min increments
A semi demountable quartz ICP torch with separate injector shall be provided for maximum flexibility.
6 | Plasma Viewing | The torch orientation should be horizontal and plasma shall be viewed axially as well as radially. Auxiliary optics shall be available to provide a radial plasma view.
The viewing height of the plasma must be adjustable under computer control to enable optimisation whilst the plasma is lit, in complete safety to the operator.
The instrument shall have the possibility to automatically switch between axial and radial view (dual view) during an analysis.
7 | Resolution | The instrument shall have the capability to perform measurements in routine operation with a Minimum optical resolution of 0.009 nanometer or better.
8 | Wavelength Range | The instrument shall be able to operate over the minimum range of 166 to 840 nm.
9 | Detector | The instrument shall be equipped with solid state CCD/CID Detector,
The Detector shall be maintained at low temperature and should able to attain the required temp in 10 mins or equivalent and background noise resulting in enhanced detection limits.
10 | Detector Operating Modes. | The detector must be capable of measuring all the analytes in the single analysis with wide linear dynamic range.
11 | Software | The instrument system software shall be based on the Windows operating system. The software shall provide full control of all instrument functions including plasma ignition, gas flows, viewing position, and monitoring of safety interlocks.
Specification
- Full instrument control with the ability to display parameters and instrument response in real time.
- Fully automated system start up and shutdown via the instrument control window.
- Real time graphics for instrument setup and monitoring.
- All set up parameters including plasma position to be stored on disk.
- Internal diagnostics including error checking and complete fault log.
- Automatic identification of possible spectral interferences when selecting wavelengths for analysis. Search mode for identification of unknown wavelengths should be available
- Automatic identification of possible spectral interferences when selecting wavelengths for analysis. Search mode for identification of unknown wavelengths etc should be available
12 | Microwave Digestion System | Vendor should provide Microwave Digestion System for 8 or more reaction vessels simultounously with individual pressure active control on reference vesel postions
Applications : For acid digestion applications of samples like plastics, water, wastewater, soil, sludge, environmental samples, natural products, plant materials, etc. for further analysis with AAS, ICP – OES or ICP – MS techniques.
Microwave Cavity : The cavity chamber must be made of corrosion resistant Stainless steel with multilayer fluoropolymer coating for physical protection as well as chemical resistance and should be less than 8 ltrs
Microwave Power : Delivered power: 1000 watts or more using 1 or 2 magnetrons
Controller : In-built color touchscreen display with real-time graphical representation of reaction parameters and display of internal temperature of every individual reaction. External/ detachable controllers not acceptable.
Rotor: System should be offered with rotor of minimum 12 position or more with 12 vessels
Rotor Stability : Rotor should be made of lightweight AL Material for high pressure strength up to 150 bar or more and to ensure ultrafast cooling.
Vessel material: PTFE- TFM Vessel volume: 50 ml or more Minimum filling volume : 3 ml or less Max operating temperature: 250 ºC or more Max operating pressure: 45 bar or more Maximum temperature: 300 ºC Maximum vessel pressure : 80 bar or more Sample weight: Minimum 1.5 gms or higher per vessels Vessel safety: Closure of the vessel must be possible by hand and overpressure release mechanism of the vessel must be controlled by means of metal springs for precise opening pressure independent of the reaction temperature or sample weight. Venting with polymer/ plastic material springs not acceptable due to dependence on internal temperature and sample weight.
Temperature calibrator: A temperature calibrator for accurate temperature measurement must be supplied along with the instrument with minimum 3 years validity
Cooling: Instrument should have built in exhaust for cooling the vessels inside the oven from 180 ºC to 70ºC in less than 15 min. Documentary evidence to be provided.
Micro sample facility: Instrument should be offered with glass inserts minimum 100 number for carrying out digestion with 1ml acid and for micro sample. Minimum three years from date of installation
13 | Chiller | Suitable Chiller/re-circulator to be provided
14 | Hydride Vapor Generator | System should be offered with Hydride/vapor generation kit
15 | Warranty | Minimum Three year from date of installation
16 | Application Note | Published application note for detection limits data should be attached.
17 | Other Accessories to be quoted and supplied along with machine / equipment | Bidder should provide 1.Argon gas cylinders (01 no`s) with double stage regulator, Nitrogen gas cylinders (02 no`s) with double stage regulator , 2.Gas Purification System, Manifold for 2 cylinders spares and consumables for 2 year operation as required 3. system with Latest PC – i7 (Dell/HP//Compaq) and Printer & Suitable online 15 KVA UPS with 60 minutes power backup to run the complete ICP system including chiller.
18 | Calibration standards | 1000 ppm of Each As, Hg, & Au Single Element Standard 100ml quantity
Limiting Oxygen Index Tester
Sl. No. | Specification | Range / Value
1 | Make / Model | Bidder to specify
2 | Purpose | To measure the minimum concentration of oxygen that will support combustion of a polymer
3 | Applicable standard | ASTM D2863, ISO 4589-1, ASTME-2931, IS-13501 and as per UL specifications
4 | Digital Read out for oxygen concentration | ±0.1%
5 | Test Chimmney | Made of heat resistant glass tube of inside diameter of 75 – 100 mm, and 400 – 500 mm height.
The bottom of the chimney or the base to which the tube is attachedshall contain non-combustible material to mix and distribute evenly the gas mixture entering at this base. Glass beads of size 3 to5 mm in diameter in a bed 80 to 100 mm deep have been found suitable. The chimney shall be mounted securely on the base toprevent air leak.
A lid of 40 mm diameter to the glass tube at the level atleast 10 mm above the top of the cylindrical chimney.
A gauze / wire screen making partition between combustible medium and tube to catch the falling fragments and keep the column clean.
6 | Specimen holder | – Small holding device that supports the specimen at its base and hold it vertically in thecenter of the chimney to be provided (For self-supporting specimen).
For samples which are flexible the specimen shall be supported by both vertical edges in aframe with referencemarks at 20 and 100 mm below the top of the frame. Theprofile of the holder and its support shall be smooth tominimize induction of turbulence in the rising flow of gas.
7 | Gas measuring and controlling devices | Capable of controlling the gas velocitythrough the chimney is 40±2 mm/s Calibrated pressure regulators and pressure gauges to be provided on the individual gas supply lines Needle valves and calibrated flow meters to be provided individualy for each gas lines
8 | Timer | capable of measuring time with an accuracy of ± 0.5 s.
Maximum: 600 sec. Digital display to be provided
9 | Flame Ignitor | Suitable flame ignitor having adjustable LPG fuel supply to be provided
9 | Features | Paramagnetic Oxygen Cell for assessing accurate oxygen (< 0.1%) levels
Compact unit for efficient use inside a laboratory hood, with ventilation
Digital selection of oxygen concentration by push-button control
Adjustment of oxygen and nitrogen flow shall be automatic to achieve the desired oxygen concentration
Digital Display: Instrument status, oxygen concentration, flow rate, chimney temperature, elapsed time
Shortened gas path for rapid response
Uniform combustion atmosphere
Dual inlet gas pressure gauges, 0-100 psi
Ignition wand with variable gas control valve
Smoke density measurement system with chart recorder
Ventilation hood compatible with the instrument to be supplied
10 | Additional Requirements | Transparent radiant heated test column
Air pump to conserve oxygen and nitrogen supply during standby period
All necessary control valves, flowmeters and filters shall be built in and the instrument making it easy to use
Suitable for Testing of Polymers, Rubbers, Fibres, Films, FRP products, Composites & allied Products
11 | Other Mandatory Items | Hard copies of Operational & Service Manual - 01 Set
Required filled gas Cylinders (Nitrogen & Oxygen) with best quality regulator
Calibration Certificate of supplied gas and flow meters
Calibration certificate traceable to NIST should be provided
Machine should come with all other essential accessories & spares required for installation, commissioning & operation.Bidder should specify and quote for any other acessories required / availble for better usage of machine.
12 | Scope of Supply / Bill of Material | Bidder should submit Scope of Supply / Bill of Material with make model of each items.
Hard copies of operational & service manual - 01 set
Two Roll Mill
Sl. No. | Specification | Range / Value
1 | Roll Diameter | mm | Minimum 150
2 | Roll Width | mm | Minimum 350
3 | Heating Power/Roll | kW | 2.5 kW- Electrical Heating
4 | Roll Speed | RPM | Variable speed Drive 2-25
5 | Nip-Gap | mm | 0.1-10 or equivalent
6 | Roll Temperature | °C | Upto 300 °C with an accuracy of ± 2 °C
7 | Accessories -Bidders should quote and supply along with the machine |  | Brass Spatula (2 Nos), Brass roll cleaning accessories,
Other necessary accessories
8 | Features |  | Ø Uniform temperature distribution across rollers
Ø Rapid cooling of rollers by water circulation
Ø Any polymer like LDPE, HDPE, PP, PVC, PS, ABS, HIPS, Natural Rubber, EVA etc. can be tested
Ø Individual adjustable roller speed with digital RPM display
Ø Incorporated with additional safety features like Knee operated reversible emergency switch at both the sides, hand operated roller safety switch and emergency push button on panel
Ø Digital nip gap displays for both rollers
Ø Control options available like HMI touch screen with PLC control and digital display by PID controls
9 | Mandatory Items |  | 1.Hard copies of operational & Service manual -01 set
2. Machine should come with all other essential accessories & spares required for installation, commissioning & Operation.
10 | Terms & Conditions |  | 1. The Supplier shall be responsible for carrying out the Installation and Commissioning at customer site.
2. Service Support: reporting time within 24 hours after officially request for service to be supplied
3. Installation & Demonstration: in the scope of Seller.
4. Loading & Unloading: in the scope of Seller.
Xenon Arc Weather Meter
Sl. No | Specification
1 | Light source - Air cooled Xenon Arc Lamps
2 | Sample compartment – Approx. 3000 Sq. cm or more
3 | Microprocessor control required
4 | Touch screen/key pad
5 | Irradiance control required
6 | Irradiance Sensors at 340nm, 420 nm & TUV range
7 | Irradiance Calibration device for quick, easy, error free and automatic calibration
8 | Optical Filter system to simulate outdoor / indoor testing
9 | Filters should be of non-ageing type to avoid frequent change
10 | Black Panel Temperature (BPT) Control
11 | Black Standard Temperature (BST) Control
12 | Chamber Air Temperature Control
13 | Simultaneous control of BST & CAT required
14 | Relative Humidity Control required
15 | Provision of water spray front & back required
16 | Data logging interface along with software
17 | Capability of mounting and testing fabric specimens/Plastic Test specimens
18 | Compliance to various international standards (ISO, ASTM and other standards) such as SAE - 2412, ASTM G155, ASTM G 151, ISO 4892-2, ASTM D 4459, ASTM D 5071, IS 15827
(B) Accessories to be quoted and supplied along with machine / equipments:
1 | Specimen holder set for Xenon test chamber.
2 | One Set of Xenon Arc Lamps for Xenon test chamber
3 | Other types of Filters (Window glass, UV/Extended UV) for Xenon test chamber
4 | Suitable Radiometer with NIST Traceable calibration certificate for optional filters for Xenon test chamber
5 | Black Panel Calibration Thermometer for Xenon test chamber
6 | Provision for back spray and dual spray in xenon test chamber