## Sheet1
GOVERNMENT OF INDIA
Central Institute of Petrochemicals Engineering & Technology
Department of Chemicals & Petrochemicals, Ministry of Chemicals & Fertilizers
No. 437/A, Hebbal Industrial Area, Mysuru – 570 016
Technical Bid
TENDER | Tender No.: CIPET:CSTS/MYS/Testing-MFI/21-22/02 Dated: February 18, 2022
NO
ITEM DESCRIPTION | Melto Flow Index Tester
Sl. No. | PARTICULARS |  | Quoted By
1 | Name of the Bidder
2 | Postal address
3 | Telephone No. with STD code
4 | Alternative Phone No. with STD Code.
5 | Name of Contact person
6 | Mobile No.
7 | e-mail ID
BELOW GIVEN DOCUMENTS TO BE SCANNED AND UPLOADED IN THE
8 | WEBSITE www.tenderwizard.com/CIPET WITHIN THE PERIOD OF SUBMISSION. |  | YES
i) Certificate of Registration for GST
ii) Certificate of Registration for NSIC / MSME / Startup / Others
iii) Auditor Certificate confirming the Turnover for the past three years 2018-19, 2019-20,2020-21 (Separate One page Auditor Certificate confirming the turnover)
9 | iv) Authorization certificate from OEM ( If not a Manufacturer for the machine / equipment) - (Section - V)
v) Bid Form (Section V)
Commericial Terms for Quoted items (Please Provide Commercial terms and conditions in the below form)
i) Minimum Required Validity of offer is 90 Days from the date of Price Bid Opening
ii) Price Shall be firm throughout the currency of contract, in the event of placement of purchase order.
iii) Terms of Payment:- Please refer Tender / Bid Document Clause nos. Accordingly payment will be released
iv) Price Term only in INR (Free Delivery / Safe Door Delivery / F O R)
v) Period required for Supply, Installation, Commissioning and Training (in days)
vi) Please Indicate the Period of Guaranty / Warranty (Minimum TWO YEARS
(2) is Mandatory) ( No. of years)
QUESTIONNAIRE TO BE FILLED BY BIDDER IN AND SENT ALONG WITH OFFER DULY SIGNED
i) Whether All Documents Relating to tender Viewed
ii) BOQ / Packing List (each and every part of the equipment & standard accessories to run the machine) - as per Tender Document
iii) Make and Model along with specification.
iv) Pre-installation requirement if any to run the equipment / machine
v) Technical Catalogue / Brochure’s / Leaflet clearly mentioning the Model number, name of equipment and technical details
vi) Proforma for Deviation Statement
11 | vii) Past three years supply details for similar machines (order copy)
viii) Performa for performance statement along with the certificate for similar machines from the client/user organization.
ix) Enter the Country of Origin of the Equipment
Make
Model
1 | Make / Model | Bidder to specify
Temperature Range | Ambient to 400⁰C. Microprocessor based PID Controller, program control and display of test results. Temperature control with two heater elements, resolution 0.01 °C 5 temperature calibration sets, each with their respective parameter files
Accuracy | ±0.1⁰C or better
Timer | Inbuilt Digital timer, with a range of 999.9 sec with buzzer indication facility, resolution better than 0.01 s
Test Chamber | Electrically heated with PID Based Digital Temperature Control. Corrosion resistance piston and barrel
Test Weights | All Weights shall be supplied as per ASTM D 1238
Die | Material: Tungsten carbide 8.0 mm ± 0.025 mm Length 2.095 ± 0.005 mm bore & as per the standards The bore of the orifice and its finish are critical. It shall have no visible drill or other tool marks and no detectable eccentricity. The bore of the orifice shall be manufactured by techniques known to produce finishes approximately 12 rms or better in accordance with ANSI B46.1.
Barrel / Cylinder | Requirement shall fulfill to all the mentioned standards
Conforms to | ASTM D 1238 (Procedures A, B, C), ISO 1133-2, IS 2530, ASTM D 3364, DIN 53735, BS 2782, IS 13360 : Part 4 : Sec 1 : Subsec 1 : 2000
Features | • Standard tungsten carbide die • Quick-release slide system for the die • Guided piston design for accurate positioning into the barrel • Integrated support for accessories and consumables • High-Precision Encoder to measure MVR, up to 40 data points acquisition for a single test • Corrosion-resistant barrel, pistons, and dies for chemically- aggressive materials, Motorized automatic Melt Cutting Device • Die Plugging Device to prevent material flowing during pre-heating • Wide range of test masses and dies for all materials and testing standards • High resolution position transducer to measure volume output, resolution 0.025 mm/impulse • Half die for procedure C and equipment compatible to perform Procedure C as per ASTM D 1238 • Enable to start the measurement automatically after the pre-melt time. • Able to measure both Melt Flow Rate & Melt Volume Rate • Suitable software with PC Control and advanced data analysis
Extra Accessories | • Calibrated Go-No-go Gauge shall be supplied • Suitable Tweezers / Spanners/ other required tools to be supplied • 02 Nos Spare Die shall be provided
Calibration and other Requirement | Machine must be compliance with international safety standards in all safety aspects. Maintenance Manuals, calibration certificates shall be provided. Vendor shall arrange installation and onsite training minimum two days for free of cost.